[image: image1.emf]
PAGE  
[image: image1.emf]

KingFisher® Flex process with 24 magnet head and MagAttract® RNA Cell Mini M48 Kit (Qiagen)

Sample process

1. Disrupt the cell pellet of max. 5 x 106 cultured cells by adding 2000 µl of RLT buffer on the cells and homogenaize according to the kit manual (MagAttract® RNA M48 Handbook, Qiagen, catalog no. 958236)
2. Fill the plates according to the Table 1.

	Plate number
	Plate type
	Plate name
	Content
	Sample/reagent volume

	1
	KingFisher Flex 24 deep well plate
	Cell Lysate
	Cell Lysate
	2000 µl

	
	
	
	Ethanol
	1500 µl

	
	
	
	MagAttract Suspension E

	400 µl

	2
	KingFisher Flex 24 deep well plate
	MW Wash
	Buffer MW
	3000 µl

	3
	KingFisher Flex 24 deep well plate
	Digestion
	Buffer RDD
	1000 µl

	
	
	
	Dnase
	50 µl

	Dispense step, add...

	3
	
	
	Buffer RPE
	3500 µl

	4
	KingFisher Flex 24 deep well plate
	RPE Wash1
	Buffer RPE
	3000 µl

	5
	KingFisher Flex 24 deep well plate
	RPE Wash2
	Buffer RPE
	3000 µl

	6
	KingFisher Flex 24 deep well plate
	Elution
	Buffer ME
	500 µl


Table 1. Filling the plates

3. Start the “Qiagen MagAttract RNA Cell Flex24” protocol using arrow keys and START button. You can also run the protocol using a computer, for more details see BindIt software user manual . 

4. Load the plates according to the protocol request and press START after every plate to confirm the action. Use KingFisher Flex 24 deep well tip comb and plate for tip loading.
Note! Confirm that the plates are placed in correct orientation: A1 well to be pointed to upper right corner of the plate holder in turntable. A1 row of the plate is then always located in the inner circle of the turntable. 
5. The purification protocol will start when the last plate is loaded and START button is pressed.

6. Dispense step: After digestion add 3500 μl of Buffer RPE to plate 3.
7. After the purification process is completed the plates are removed according to instructions shown in the instrument screen. Press START after each plate removal to confirm the action.
8. When the last plate is removed text End of run will appear. Press STOP to complete the run.
For more information: www.thermo.com/kingfisher and www.qiagen.com





www.thermo.com


[image: image2.png]ELECTRON CORPORATION



