[image: image1.emf]
PAGE  
[image: image1.emf]

Thermo Scientific KingFisher® Flex process with 24 magnet head and ChargeSwitch® Total RNA Cell kit (Invitrogen)
Sample process

1. Prepare the cell lysate of max. 5 x 106 cultured cells according to the sample preparation of the kit manual (ChargeSwitch® Total RNA Cell Kit, Invitrogen, catalog no. CS14010). Use 2500 µl of Lysis mix solution (Lysis buffer L16, DTT, Proteinase K). 
2. Fill the plates according to the Table 1.

	Plate number
	Plate type
	Plate name
	Content
	Sample/reagent volume

	1
	KingFisher Flex 24 deep well plate
	Lysis
	Cell Lysate
	2500 µl

	
	
	
	Binding Buffer B9
	1000 µl

	
	
	
	ChargeSwitch® Magnetic Beads
	500 µl

	2
	KingFisher Flex 24 deep well plate
	Wash 1
	Wash Buffer W14
	3000 µl

	3
	KingFisher Flex 24 deep well plate
	Dnase I
	Dnase I mix
	1250 µl

	Dispense step, add...

	3
	
	
	Binding Buffer B9
	400 µl

	4
	KingFisher Flex 24 deep well plate
	Wash 2
	Wash Buffer W13
	4000 µl

	5
	KingFisher Flex 24 deep well plate
	Wash 3
	Wash Buffer W14
	3000 µl

	6
	KingFisher Flex 24 deep well plate
	Elution
	Elution Buffer E7
	500 µl


Table 1. Filling the plates

3. Start the “Invitrogen Total RNA Flex24” protocol using arrow keys and START button. You can also run the protocol using a computer, for more details see BindIt software user manual . 

4. Load the plates according to the protocol request and press START after every plate to confirm the action. Use KingFisher Flex 24 deep well tip comb and plate for tip loading.
Note! Confirm that the plates are placed in correct orientation: A1 well to be pointed to upper right corner of the plate holder in turntable. A1 row of the plate is then always located in the inner circle of the turntable. 
5. The purification protocol will start when the last plate is loaded and START button is pressed.

6. Dispense step: After DNase I digestion add 400 μl of Binding Buffer B9 to plate 3.
7. After the purification process is completed the plates are removed according to instructions shown in the instrument screen. Press START after each plate removal to confirm the action.
8. When the last plate is removed text End of run will appear. Press STOP to complete the run.
For more information: www.thermo.com/kingfisher and www.invitrogen.com





www.thermo.com


[image: image2.png]ELECTRON CORPORATION



