KingFisher® Flex process with 96 deep well magnet head and InviMag® Plant DNA Kit (Invitek)

Sample process

- Homogenize and lyse the sample according to the InviMag® Plant DNA Kit Manual, Invitek, catalog no. 74373001

- Fill the plates according to the Table 1. More information about the reagents in the InviMag® Plant DNA Kit Manual

| Plate number | Plate type | Plate name | Content | Sample/reagent volume |
| --- | --- | --- | --- | --- |
| 1 | Microtiter deep well 96 plate | Lysate | Binding Buffer | 200 µl |
|  |  |  | Sample lysate | 200 µl |
|  |  |  | SNAP Solution | 20 µl |
| 2 | Microtiter deep well 96 plate | Washing Buffer I | Washing Buffer I | 800 µl |
| 3 | Microtiter deep well 96 plate | Washing Buffer II | Washing Buffer II | 800 µl |
| 4 | Microtiter deep well 96 plate | Washing Buffer II_2 | Washing Buffer II | 800 µl |
| 5 | KingFisher 96 plate | Elution | Elution Buffer | 100 µl |

Table 1. Filling the plates

- Combine the 96 deep well tip comb and the KingFisher 96 plate. For more detailed instructions, see KingFisher Flex User manual

- Start the “InviMag PlantDNA_Flex96” protocol using arrow keys and START button. You can also run the protocol using a computer, for more details see BindIt software user manual

- Load the plates according to the protocol request and press START after every plate to confirm the action.

Note! Confirm that the plates are placed in correct orientation: A1 well to be pointed to upper right corner of the plate holder in turntable. A1 row of the plate is then always located in the inner circle of the turntable

- The purification protocol will start when the last plate is loaded and START button is pressed

- After the purification process is completed the plates are removed according to instructions shown in the instrument screen. Press START after each plate removal to confirm the action

- When the last plate is removed text End of run will appear. Press STOP to complete the run

For more information: www.thermo.com/kingfisher and www.invitek.eu