KingFisher®  Flex process with 96 deep well magnet head and sbeadex®  Plant DNA Kit (AGOWA)

Sample process

- Homogenize and lyse the sample according to the sbeadex® Plant DNA Kit instructions, AGOWA, catalog no. 41220

- Fill the plates according to the Table 1. More information about the reagents in the sbeadex® Plant DNA Kit instructions

| Plate number | Plate type | Plate name | Content | Sample/reagent volume |
| --- | --- | --- | --- | --- |
| 1 | Microtiter deep well 96 plate | Lysate | Binding Buffer PN | 360 µl |
|  |  |  | Sample Lysate | 150 µl |
|  |  |  | Particles suspension PN | 30 µl |
| 2 | Microtiter deep well 96 plate | Wash 1 | Wash Buffer PN1 | 500 µl |
| 3 | Microtiter deep well 96 plate | Wash 2 | Wash Buffer PN2 | 500 µl |
| 4 | Microtiter deep well 96 plate | Wash 3 | Aqua | 500 µl |
| 5 | KingFisher 96 plate | Elution | Elution Buffer PN | 150 µl |

Table 1. Filling the plates

- Combine the 96 deep well tip comb and the KingFisher 96 plate. For more detailed instructions, see KingFisher Flex User manual

- Start the “AGOWA PlantDNA Flex96” protocol using arrow keys and START button. You can also run the protocol using a computer, for more details see BindIt software user manual

- Load the plates according to the protocol request and press START after every plate to confirm the action

Note! Confirm that the plates are placed in correct orientation: A1 well to be pointed to upper right corner of the plate holder in turntable. A1 row of the plate is then always located in the inner circle of the turntable.

- The purification protocol will start when the last plate is loaded and START button is pressed.

- After the purification process is completed the plates are removed according to instructions shown in the instrument screen. Press START after each plate removal to confirm the action

- When the last plate is removed text End of run will appear. Press STOP to complete the run

For more information: www.thermo.com/kingfisher and www.agowa.de