KingFisher®  Flex process with 96 deep well magnet head and NucleoMag™ 96 Virus Kit (Macherey-Nagel)

Sample process

- Fill the plates according to the Table 1. More information about the reagents in the NucleoMag™ 96 Virus Kit instructions, Macherey-Nagel, catalog no. 744 800

| Plate number | Plate type | Plate name | Content | Sample/reagent volume |
| --- | --- | --- | --- | --- |
| 1 | Microtiter deep well 96 plate | Lysis | Lysis Buffer MV1 | 200 µl |
|  |  |  | Sample | 200 µl |
|  |  |  | Proteinase K | 10 µl |
|  |  |  | Carrier RNA | 4 µl |
| Dispense step, add... |
| 1 |  |  | Binding Buffer MV2 | 550 µl |
|  |  |  | NucleoMag V-Beads | 30 µl |
| 2 | Microtiter deep well 96 plate | Wash_MV3 | Wash Buffer MV3 | 500 µl |
| 3 | Microtiter deep well 96 plate | Wash_MV4 | Wash Buffer MV4 | 500 µl |
| 4 | Microtiter deep well 96 plate | Wash_MV5 | Wash Buffer MV5 | 550 µl |
| 6 | KingFisher 96 plate | Elution | Elution Buffer MV6 | 100 µl |

Table 1. Filling the plates

- Combine the 96 deep well tip comb and the KingFisher 96 plate. For more detailed instructions, see KingFisher Flex User manual

- Start the “NucleoMag Virus Flex96” protocol using arrow keys and START button. You can also run the protocol using a computer, for more details see BindIt software user manual

- Load the plates according to the protocol request and press START after every plate to confirm the action

Note! Confirm that the plates are placed in correct orientation: A1 well to be pointed to upper right corner of the plate holder in turntable. A1 row of the plate is then always located in the inner circle of the turntable

- The purification protocol will start when the last plate is loaded and START button is pressed.

- Dispense step: After lysis add 30 μl of resuspended NucleoMag V-Beads and 550 μl of Binding Buffer MV2 to plate 1. Beads and MV2 may be premixed before addition to plate 1

- After the purification process is completed the plates are removed according to instructions shown in the instrument screen. Press START after each plate removal to confirm the action

- When the last plate is removed text End of run will appear. Press STOP to complete the run

For more information: www.thermo.com/kingfisher and www.mn-net.com